## C3D Scale
C3D Analog Scale Calculations
These sheets enable the user to calculate the correct C3D scale factors
required by software using C3D files. See the help page for details.
Manufacturer specific sheets for:
| Advanced Mechanical Technology Inc. |  | Force Plate
| Bertec |  | Force Plate
| Kistler |  | Force Plate
| Motion Lab Systems, Inc. |  | Electromyography
| Sensotec |  | Load Cell
If you find any errors in these calculations please send details to support@motion-labs.com
Please visit the C3D.ORG web site for complete information about the C3D file format.
Contact your equipment manufacturer for information about C3D scale values.
| Motion Lab Systems, Inc | Tel: | +1 225 272-7364
15045 Old Hammond Highway
| Baton Rouge, LA 70816 USA | Email: | support@motion-labs.com
|  |  | Revised : | April 22, 2020

## Help
How to use these spreadsheets
These spreadsheets enable you to calculate the correct values for the C3D ANALOG:SCALE parameters. This value is stored within the C3D file when it is created and is used to scale the stored digital samples into "real-world" units like volts, degrees, and force. The spreadsheets require information about the data collection and/or measurement system in order to calculate these values - refer to the C3D format manual supplied with your equipment or available on the internet at www.c3d.org for specific details and worked examples of the formulas used in these sheets.
Enter the information about your data collection environment - the values shown in red describe your ADC card setup, while blue values relate to the calibration factors (if any) for the signal source. Make sure that you use the correct ANALOG:GEN_SCALE factor for your system - contact your data collection system manufacturer if you have any questions. The appropriate ANALOG:SCALE factor for your device or analog channel will be shown in green - this value is calculated from the information that you provide.
The worksheets require the following information:
|  | BLUE | Manufacturers calibration or cross-talk information.
|  | RED | Data collection setup parameters.
The worksheets generate the following information based on the input values.
|  | GREEN | ANALOG:SCALE parameter values.

## Parameters
How to calculate C3D analog parameters
Enter your ADC setting (shown in red) to see the optimum values shown in green.
| ADC input range | 20 | Volts
| Overall ADC gain (x1) | 1 |  | Overall ADC input range is ± | 10 | Volts
| ADC resolution | 12 | bits | Total of | 4096 | bits of resolution
| ADC channel gain | 1 | Individual ADC channel gain (x1, x2, x4, or x8)
|  |  |  | The ADC Input range is ± | 10 | Volts
| One bit represents | 0.0048828 | Volts, i.e. | 4.8828 | mV.
|  |  | When the ANALOG:GEN_SCALE parameter is | 0.0048828
|  |  | The individual channel ANALOG:SCALE parameter is | 1 | to scale the ADC samples in volts.
|  |  | The default ANALOG:OFFSET parameter is | 2047 | for the ADC described above.
Note that a 12-bit ADC has 4096 bits of resolution and generates a binary value where 000000000000 is the minimum sample value and 111111111111 is the maximum value. The output is a zero based count, not a one based count. As a result the maximum value that a 12-bit ADC can generate is 4095, not 4096, so the center of the signal range is 2047, not 2048.
[5 empty] | 0000111111111111 |  | ADC resolution mask
| For example | 4096 | translated to 16 bit binary is |  | 0001000000000000 |  | ADC range ERROR
Note on C3: This is the ADC card voltage range, normally quoted as "±10 Volts" which which means that the range of the ADC is 20 Volts from +10V to -10V.
Note on C4: Used by older ADCs which have a single gain control that affects all sampled channels. Most modern ADCs use individual channel gain settings. The default value is 1 if there is no gain option.
Note on C5: The number of bits of resolution available to the ADC device - usually 12, 14, or 16 bits.
Note on C7: The individual ADC channel gain affects the individual ADC channel input range. When a ±10V ADC channel gain is "2" then the ADC channel input range is ±5V range. This value is not recorded in the C3D file but affects each individual ANALOG:SCALE calculation.
Note on E12: Enter your GEN_SCALE parameter value here. Common C3D values are normally 1.0 or 0.004883. The ANALOG:GEN_SCALE is always a positive value. Individual ANALOG:SCALE parameters may be either positive or negative. A negative value will invert the scaled data values

## Calculations
C3D file ADC analog scale calculations
Enter your ADC setting (shown in red) to see the optimum values shown in green.
| ADC input range | 20 | Volts
| Overall ADC input range is ± | 10 | Volts
| ADC resolution | 16 | bits | 65536 | bits of resolution
| Individual ADC channel gain | 2 | (typically x1, x2, x4, or x8)
| Thus the signal range is ± | 5 | Volts. This is determined by the gain applied to the analog channel.
| Therefore 1 bit represents | 0.00015259 | Volts, i.e. | 0.15259 | mV.
|  | When the ANALOG:GEN_SCALE parameter is | 0.00030518 |  | the default ANALOG:SCALE is | 0.5
[5 empty] | The ADC resolution defines the default ANALOG:OFFSET parameter as | 32767
The calculation below translates a floating-point or integer C3D file analog data value into volts
|  | When ANALOG:GEN_SCALE= | 0.0003052 |  |  | 0
|  | ANALOG:SCALE = | 0.5 |  |  | 0
|  | ANALOG:OFFSET = | 32767
|  | A C3D analog data value of | 32768 | represents | 0.0001526 | Volts
Note on E5: The ADC samples are binary values and always return values from 0 to the maximum ADC resolution which is the number of ADC bits to the power of two, minus one (because 0 is a valid sample).

## MA-400
MA300/400 EMG scale calculation for volts at skin surface
| ADC input range | 20 | Volts (+10 volts to -10 volts)
| ADC channel gain (x1) | 2 |  | ADC input range is ± | 5 | Volts
| ADC resolution | 12 | bits, therefore | 4096 | bits ADC resolution.
| One bit represents | 0.0024414 | Volts
| ANALOG:GEN_SCALE | 0.0024414
MA400 EMG Gain switch settings |  |  |  | MA400 output range is ± | 5 | Volts
| Switch | Total Gain | Max EMG input (mV) |  | ANALOG:SCALE
| 0 | 350 | 25.974 |  | 0.0028572
| 1 | 2000 | 4.5455 |  | 0.0005
| 2 | 4000 | 2.2727 |  | 0.00025
| 3 | 5700 | 1.5949 |  | 0.00017544
| 4 | 8000 | 1.1364 |  | 0.000125
| 5 | 9500 | 0.95694 |  | 0.00010526
| 6 | 11500 | 0.79051 |  | 8.6957e-05
| 7 | 13200 | 0.68871 |  | 7.5758e-05
| 8 | 16600 | 0.54765 |  | 6.0241e-05
| 9 | 18000 | 0.50505 |  | 5.5556e-05
Note on C3: ADC card voltage range - normally quoted as "+/- 10 Volts" - which means that the total range of the ADC is 20 Volts.
Note on C4: ADC channel gain - set to "2" for a +/-5V range with most ADC cards.
Note on C5: The number of bits of resolution available to the ADC device - usually 12, 14, or 16.
Note on C7: The C3D ANALOG:GEN_SCALE parameter is applied to all analog channel calculations and should be set to the value of 1 bit.
Note on D10: This is the maximum input level reduced by 10% to avoid clipping the input signal.

## Load Cell
Load Cell Scale Calculation
| ADC input range | 20 | Volts (+10 volts to -10 volts)
| ADC channel gain (x1) | 8 |  | Signal range is +/- | 1.25 | Volts
| ADC resolution | 12 | bits |  | 4096 | bits
| One bit represents | 0.00061035 | Volts
Quick Calculator
| ANALOG:GEN_SCALE | 1 |  | Input | 0.02 | Volts
|  |  |  | ADC count | 2081 | Bits
|  |  |  | Output | 50.354 | (units)
This example uses a Sensotec Model 31 Precision Minature Load
Cell with an output of 2.395mV per volt of excitation at full load.
| Maximum Load | 50 | (units)
| Excitation Voltage | 10 | Volts
| Calibration Factor | 0.002 | Volts out per volt of excitation at Full Load
| External Gain | 0 | Enter gain of signal conditioning unit (if any)
|  | 0.0004 | Volts output per (unit) of load.
| ANALOG:SCALE | 1.5259
Note on C3: ADC card voltage range - normally quoted as "+/- 10 Volts" - which means that the total range of the ADC is 20 Volts.
Note on C4: ADC channel gain
Note on C5: The number of bits of resolution available to the ADC device - usually 12 or 16.
Note on C8: Enter your GEN_SCALE value here. Common values are either 1.0 or 0.004883

## AMTI-2
AMTI Force Plate - TYPE-2 scales
| ADC input range | 20 | Volts (+10 volts to -10 volts)
| ADC gain (x1) | 1 |  | Signal range is +/- | 10 | Volts
| ADC resolution | 12 | bits |  | 4096 | bits
| One bit represents | 0.0048828 | Volts per bit.
| ANALOG:GEN_SCALE | 1 |  | AMTI filters are set at either 10Hz or
| ANALOG:OFFSET | 2047 |  | 1020Hz by switches on the amplifier.
| Excitation Voltage | 10 | Volts
| Cable length (in feet) | 30 | See AMTI Technical note TN191-01
Force Plate | S Matrix | Excitation
Sensitivity | uV/Vo/N | Factor V/Ft | Gain | Chan Gain | ANALOG:SCALE
S [Fx,Fx'] | 0.655 | 0.0003708 | 4000 | 1 | -0.18657 | Newtons per Volt
S [Fy,Fy'] | 0.65 | 0.0003708 | 4000 | 1 | -0.18801 | Newtons per Volt
S [Fz,Fz'] | 0.167 | 0.0001852 | 4000 | 1 | -0.73137 | Newtons per Volt
S [Mx,Mx'] | 1.348 | 0.0003708 | 4000 | 1 | -90.6575 | Newton-mm/Volt
S [My,My'] | 1.352 | 0.0003708 | 4000 | 1 | -90.3892 | Newton-mm/Volt
S [Mz,Mz'] | 2.552 | 0.0003708 | 4000 | 1 | -47.8865 | Newton-mm/Volt
Note on C3: ADC card voltage range - normally quoted as "+/- 10 Volts" - which means that the total range of the ADC is 20 Volts.
Note on C4: Overall ADC gain: Used by Vicon-VX but not the 370/512 Vicon systems which use individual channel gain settings.
Note on C5: The number of bits of resolution available to the ADC device - usually 12 or 16.
Note on C8: Enter your GEN_SCALE value here. Common values are either 1.0 or 0.004883
Note on C11: Force Plate Excitation voltage: This is set inside the AMTI amplifier box.
Note on C12: Enter the length of the cable between the AMTI amp and the Force Plate. See note TN191-01 from AMTI for cable length information.
Note on B14: Older AMTI force plates should use the C' matrix (SI units) data provided with the force plate. Newer AMTI force plates should use the S matrix (SI units) data.
Note on D15: AMTI Channel gain - this is set by DIP switches on the AMTI amplifier.
Note on E15: Individual ADC channel gain as applied by the Analog Data Collection card.

## AMTI-2a
AMTI Force Plate - Calculate matrix values from ANALOG:SCALE values
| ADC input range | 20 | Volts (+10 volts to -10 volts)
| ADC gain (x1) | 1 |  | Signal range is +/- | 10 | Volts
| ADC resolution | 12 | bits |  | 4096 | bits
| One bit represents | 0.0048828 | Volts per bit.
| ANALOG:GEN_SCALE | 1 |  | This sheet calculates the S-Matrix values given the ANALOG:SCALE values found in the C3D file. It is provided so that users can determine the force plate assignments based on the data stored in a C3D file.
| ANALOG:OFFSET | 2047
| Excitation Voltage | 10 | Volts
| Cable length (in feet) | 30 | See AMTI Technical note TN191-01
Excitation
ANALOG:SCALE |  | Factor V/Ft | Gain | Chan Gain | S Matrix - uV/Vo/N
-0.187 |  | 0.0003708 | 4000 | 1 | 0.65351 | S [Fx,Fx']
-0.1888 |  | 0.0003708 | 4000 | 1 | 0.64728 | S [Fy,Fy']
-0.7385 |  | 0.0001852 | 4000 | 1 | 0.16539 | S [Fz,Fz']
-91.083 |  | 0.0003708 | 4000 | 1 | 1.3417 | S [Mx,Mx']
-90.101 |  | 0.0003708 | 4000 | 1 | 1.3563 | S [My,My']
-47.874 |  | 0.0003708 | 4000 | 1 | 2.5527 | S [Mz,Mz']
Note on D3: ADC card voltage range - normally quoted as "+/- 10 Volts" - which means that the total range of the ADC is 20 Volts.
Note on D4: Overall ADC gain: Used by Vicon-VX but not the 370/512 Vicon systems which use individual channel gain settings.
Note on D5: The number of bits of resolution available to the ADC device - usually 12 or 16.
Note on D8: Enter your GEN_SCALE value here. Common values are either 1.0 or 0.004883
Note on D11: Force Plate Excitation voltage: This is set inside the AMTI amplifier box.
Note on D12: Enter the length of the cable between the AMTI amp and the Force Plate. See note TN191-01 from AMTI for cable length information.
Note on E15: AMTI Channel gain - this is set by DIP switches on the AMTI amplifier.
Note on F15: Individual ADC channel gain as applied by the Analog Data Collection card.

## AMTI-4
AMTI Force Plate - TYPE-4 scales
| ADC input range | 20 | Volts (+10 volts to -10 volts)
| ADC gain (x1) | 1 |  | Signal range is +/- | 10 | Volts
| ADC resolution | 12 | bits |  | 4096 | bits
| One bit represents | 0.0048828 | Volts per bit.
| ANALOG:GEN_SCALE | 1 |  | AMTI filters are set at either 10Hz or
| ANALOG:OFFSET | 2047 |  | 1020Hz by switches on the amplifier.
| Excitation Voltage | 10 | Volts
| Cable length (in feet) | 30 | See AMTI Technical note TN191-01
Excitation
|  | Factor V/Ft | Gain | Chan Gain | ANALOG:SCALE
|  | 0.0003708 | 4000 | 1 | -0.12221 | Newtons per Volt
|  | 0.0003708 | 4000 | 1 | -0.12221 | Newtons per Volt
|  | 0.0001852 | 4000 | 1 | -0.12214 | Newtons per Volt
|  | 0.0003708 | 4000 | 1 | -0.12221 | Newtons per Volt
|  | 0.0003708 | 4000 | 1 | -0.12221 | Newtons per Volt
|  | 0.0003708 | 4000 | 1 | -0.12221 | Newtons per Volt
Note on C3: ADC card voltage range - normally quoted as "+/- 10 Volts" - which means that the total range of the ADC is 20 Volts.
Note on C4: Overall ADC gain: Used by Vicon-VX but not the 370/512 Vicon systems which use individual channel gain settings.
Note on C5: The number of bits of resolution available to the ADC device - usually 12 or 16.
Note on C8: ANALOG:GEN_SCALE value: Although commonly set to 1.0 it is recommended that the value of 0.004883 is used as some software applications are incapable of handling very small SCALE values.
Note on C11: Force Plate Excitation voltage: This is set inside the AMTI amplifier box.
Note on C12: Enter the length of the cable between the AMTI amp and the Force Plate. See note TN191-01 from AMTI for cable length information.
Note on D15: AMTI Channel gain - this is set by DIP switches on the AMTI amplifier.
Note on E15: Individual ADC channel gain as applied by the Analog Data Collection card.

## Bertec-2
Bertec Force Plate - TYPE-2 scales
| ADC input range | 20 | Volts (+10 volts to -10 volts)
| ADC gain (x1) | 1 |  | Signal range is +/- | 10 | Volts
| ADC resolution | 12 | bits |  | 4096 | bits
| One bit represents | 0.0048828 | Volts per bit.
| ANALOG:GEN_SCALE | 1
| ANALOG:OFFSET | 2047
Bertec plates have internal filters at 500Hz.
Force Plate
Sensitivity | C matrix N/V and Nm/V |  | Gain | Chan Gain | ANALOG:SCALE
S [Fx,Fx'] | 1455 |  | 10 | 1 | -0.71045 | Newtons per Volt
S [Fy,Fy'] | 1443 |  | 10 | 1 | -0.70459 | Newtons per Volt
S [Fz,Fz'] | 2887 |  | 10 | 1 | -1.4097 | Newtons per Volt
S [Mx,Mx'] | 767 |  | 20 | 1 | -187.2559 | Newton-mm/Volt
S [My,My'] | 533 |  | 20 | 1 | -130.127 | Newton-mm/Volt
S [Mz,Mz'] | 337 |  | 20 | 1 | -82.2754 | Newton-mm/Volt
Note on C3: ADC card voltage range - normally quoted as "+/- 10 Volts" - which means that the total range of the ADC is 20 Volts.
Note on C4: Overall ADC gain: Used by Vicon-VX but not the 370/512 Vicon systems which use individual channel gain settings.
Note on C5: The number of bits of resolution available to the ADC device - usually 12 or 16.
Note on C8: ANALOG:GEN_SCALE value: Although commonly set to 1.0 it is recommended that the value of 0.004883 is used as some software applications are incapable of handling very small SCALE values.
Note on B15: Values for the "C" matrix (below) are obtained from the Bertec documentation suplied with the plate.
Note on D15: This is the gain setting on the Bertec amplifier which usually cover the following ranges: x1, x2, x5, x10, x20, x50, x100
Note on E15: Gain setting of the individual ADC card channel.
Note on F15: Individual ADC scale factors.

## Kistler-3
Kistler Force Plate - TYPE-3 scales
| ADC input range | 20 | Volts (+10 volts to -10 volts)
| ADC gain (x1) | 1 |  | Signal range is +/- | 10 | Volts
| ADC resolution | 12 | bits |  | 4096 | bits
| One bit represents | 0.0048828 | Volts per bit.
| ANALOG:GEN_SCALE | 1
| ANALOG:OFFSET | 2047
Enter the Sensitivity setting from the Kistler calibration sheet and
the range as set on the front panel of the force plate amplifiers.
Channel Gain is the individual ADC channel setting.
| Sensitivity |  | Range | Ch Gain | pC/V * NpC * V/bit
| Fx (pC/N) | 7.87 | 5000 | 1 | -0.31022 | Newtons per Volt
| Fy (pC/N) | 7.85 | 5000 | 1 | -0.31101 | Newtons per Volt
| Fz (pC/N) | 3.89 | 10000 | 1 | -1.2552 | Newtons per Volt
Note on C3: ADC card voltage range - normally quoted as "+/- 10 Volts" - which means that the total range of the ADC is 20 Volts.
Note on C4: Overall ADC gain: Used by Vicon-VX but not the 370/512 Vicon systems which use individual channel gain settings.
Note on C5: The number of bits of resolution available to the ADC device - usually 12 or 16.